## Sheet1
USC Norris Healthcare Center - Medical Equipment Log
ITEM DESCRIPTION | Manufacturer | Model# | N | E | F | H | W | D | GRP | SEE DIMS ON MFGR SPEC | WEIGHT (LBS) | OP WEIGHT (LBS) | PLACEMENT/MOUNTING | ANCH | VOLTS | PHASE | AMPS | WATTS | VA | CIRCUIT BREAKER SIZE | DEDICATED | EMERGENCY | HARDWIRED | POWERCORD | NETWORK PORT | PHONE LINE | REMOTE ALARM | OTHER ELEC REQS | MED GAS | CMPRSD AIR | EXT. EXHAUST | VENTILATION | BLDG STEAM | INT STEAM GEN | COLD H20 | HOT H20 | TREATED H20 | DRAIN | BTUs | NEW FURN/INST BY | SEE COMMENTS ON CUTSHEET SUMMARY
313 - Microwave, Countertop 1.8 CUFT | GE Appliances | JEB1860DMWW | 8 | - | - | 14 | 24 | 19-1/2 | 2 |  | 43 |  | Countertop | POR | 120 | 1 | 14 | 1550 | [5 empty] | X | [15 empty] | O/O
503 - Cart, Transfer, Universal | Steris | FD61-700 | 5 | - | - | 37 | 31 | 38 |  |  | 145 |  | Mobile | POR | [25 empty] | O/O
525 - Waste Receptacle, 44 Gal, Biohazard | Rubbermaid Commercial | 2643-60 | 14 | - | - | 31-1/2 | 24 | 24 | 3 |  | 52 | 102 | Mobile | POR | [25 empty] | O/O
999-04 - Chemical Dispensing System, Acu-hold Container, 3 Pump Delivery System | Steris | PROLYSTICA ACU-HOLD | 2 | - | - | 14 | 10-1/4 | 6 | 2 | X |  |  | WallMount | VIC | 120 | 1 | [7 empty] | X | [15 empty] | O/V | X
1186 - Rack, Sterile Wrap | Intermetro Industries Corp | SWR566DC/2460CI-4 | 1 | - | - | 68 | 24 | 60 | 3 | X | 94 |  | Mobile | POR | [25 empty] | O/O
1542 - C-Locker, Wall Mount | Herman Miller | CO561FF | 11 | - | - | 78 | 23 | 20 | 1 |  |  |  | WallMount | WAL | [25 empty] | O/C
1861 - Cart, Housekeeping | Royce Rolls Ringer | F36-08E | 11 | - | - | 43 | 23 | 36 | 2 |  |  |  | Mobile | POR | [25 empty] | O/O
2142 - Hamper, Linen (for Containment of Soiled Linen) | Blickman | 2301SS | 47 | - | - | 32 | 18 | 18 | 3 |  | 25 |  | Mobile | POR | [25 empty] | O/O
3376-1 - Compressor, Air, Rotary Scroll | Steris | STS 050142 | 1 | - | - | 35 | 40 | 24 | 2 | X | 260 |  | Freestanding | VIC | 208 | 3 | 14 |  |  |  | X | X |  | X | [15 empty] | O/V | X
3845 - Waste Receptacle, General, 18 GAL, 27"HX 17"W X 20"D | Rubbermaid Commercial | 6145W | 41 | - | - | 27 | 17 | 20 |  |  |  |  | Mobile | POR | [25 empty] | O/O
6146 - Waste Receptacle, Biohazard, 23 Gal, 33in. H X 17in. W X 20in. D | Rubbermaid Commercial | 6146 BIOHAZARD | 16 | - | - | 33 | 17 | 20 | 3 |  |  |  | Mobile | POR | [25 empty] | O/O
6848 - Cart, Wire 24D 48W 69H | Intermetro Industries Corp | N556EC/2448CI-4 | 18 | - | - | 69 | 48 | 24 | 3 | X | 120 | 1020 | Mobile | POR | [25 empty] | O/O
6872 - Cart, Wire 24D 60W 69H | Intermetro Industries Corp | SUPER ERECTA N566EC/ | 36 | - | - | 69 | 60 | 24 | 3 |  | 112 | 912 | Mobile | POR | [25 empty] | O/O | X
8113 - Rack Return | Steris | RELIANCE RACK RETURN | 1 | - | - | 31 | 37 | 11 | 1 | X | 235 |  | Recessed -W | VIC | [25 empty] | O/C | X
9485 - Load_Unload Conveyor, Single, Automated | Steris | TBD | 2 | - | - | 31 | 32 | 32 | 2 | X | 94 | 229 | Freestanding | FLR | [6 empty] | X | X | X | [6 empty] | X | [7 empty] | X |  | O/V
9485-1 - Load_Unload Conveyor, Single, Manual | Steris | SCS-1L/SCS-1U | 4 | - | - | 31 | 32 | 32 | 2 | X | 94 | 229 | Freestanding | FLR | 120 | 1 | 15 | [22 empty] | O/V
9638 - Storage System, High Density, Top-Track | Intermetro Industries Corp | HIGH DENSITY TOP- TRA | 2 | - | - | - | - | - | 3 | X |  |  | Freestanding | FLR | [25 empty] | O/O
10459 - Sealer, Heat, Countertop (Temperatures Digitally Controlled) | Rennco Incorporated | LIFT SEAL LS-18D | 2 | - | - | 8 | 20 | 11 | 2 |  | 46 | 46 | Countertop |  | 120 | 1 | 5 | [6 empty] | X | [15 empty] | O/O
11831 - Reprocessor, Endoscope | Advanced Sterilization Product | EVOTECH | 2 | - | 1 | 58 | 46 | 33 | 2 | X | 500 | 578 | Freestanding | FLR | 208 | 3 |  | 9000 |  | 30 | X |  |  | X | X |  |  | X | [6 empty] | X |  | X | X | 7000 | O/C | X
16062-1 - Cart, Instrument Wash, 55-15_16in. H X 65in. W X 32.5in. D | Steris | INSTRUMENT WASH CA | 1 | - | - | 56 | 65 | 33 | 3 |  |  |  | Mobile | POR | [25 empty] | O/O
16062-2 - Cart, Container Wash, 71in. H X 30in. W X 66in. D | Steris | CONTAINER WASH CART | 2 | - | - | 70-3/4 | 30 | 66 | 3 |  |  |  | Mobile | POR | [25 empty] | O/O
16575 - Cart, Case, Closed 39H X 33W X 29D | Intermetro Industries Corp | CASE24-L6S | 25 | - | - | 39 | 33 | 29 | 3 |  | 115 |  | Mobile | POR | [25 empty] | O/O
17267 - Transporter, C-Locker | Herman Miller | CO342FF | 2 | - | - | 42 | 28 | 26 | 3 |  |  |  | Mobile | POR | [25 empty] | O/O
19494-54 - Loading Car_Transfer Carriage for 54in. Chamber Sterilizer | Steris | EVOLUTION | 4 | - | - | 58 | 25 | 78 | 3 | X |  |  | Mobile | POR | [25 empty] | O/O
19816 - Sonic Cleaner, Countertop | Steris | RELIANCE CRT5 | 1 | - | - | 13 | 13 | 21 | 2 | X | 36 |  | Countertop | CTP | 120 | 1 | 12 | [6 empty] | X | [15 empty] | O/O
19892 - Washer Disinfector | Steris | RELIANCE VISION | 2 | - | - | 81 | 42 | 32 | 1 | X | 458 | 1187 | Recessed -W | FLR | 480 | 3 | 17 |  |  | 25 |  |  | X |  | X |  |  |  |  | X | X |  | X |  | X | X | X | X | 6299 | O/V
22143 - Sterilizer, 26 x 37.5 x 54 | Steris | EVOLUTION | 2 | - | - | 78 | 38 | 76 | 1 | X | 3800 | 3800 | Freestanding | FLR | 480 | 3 | 6 |  |  | 8 | X | X | X |  | X |  |  |  |  | X |  |  | X |  | X |  | X | X | 14650 | O/V
23162 - Cabinet, Endoscope | Stanley Innerspace Corporation | 4800AA | 2 | - | - | 92 | 36 | 18 | 1 |  | 205 | 2100 | WallMount | WAL | 120 | [24 empty] | O/C
23724 - Soak Station, GUS, Countertop | PCI Medical | G32-E | 1 | - | - | 27 | 36 | 27 | 2 |  | 110 |  | Countertop | CTP | 120 | 1 | 1 |  |  |  |  | X |  | X | [15 empty] | O/O | X
24644 - Washer, Cart, Pit Mounted | Steris | VISION 1327 | 1 | - | - | 112-3/4 | 124 | 130 | 1 | X | 6367 | 7483 | Recessed -W | FLR | 480 | 3 | 14 | [5 empty] | X |  | X |  |  |  |  | X | X |  | X |  | X | X | X | X | 13800 | O/V
24683 - Changing Station, Baby, Horizontal | Koala Kare Products | KB200 | 4 | - | - | 22 | 35 | 23 | 1 |  | 28 |  | WallMount | WAL | [25 empty] | O/C
24684-1 - Cart, Scope Transport, 6-Tray | Advanced Sterilization Product | CLEANASCOPE 103600 | 1 | - | - | 36-1/2 | 20-1/2 | 23 | 3 |  |  |  | Mobile | POR | [25 empty] | O/O
24684-2 - Cart, Scope Transport, 10-Tray | Advanced Sterilization Product | CLEANASCOPE 103635 | 1 | - | - | 56-1/2 | 20-1/2 | 23 |  |  |  |  | Mobile | POR | [25 empty] | O/V
25542 - Ultrasonic Cleaner, Console | Steris | CAVIWAVE PRO CRP117 | 1 | - | - | 59 | 36 | 35 | 2 | X |  | 900 | Freestanding | VIC | 208 | 3 | 27 | [5 empty] | X | [11 empty] | X | X | X | X | 300 | O/C | X
25575 - Sterilizer, Low Temperature-1 Door | Steris | V-PRO MAX | 2 | - | - | 76 | 33 | 39 | 2 | X |  | 780 | Freestanding | FLR | 208 | 3 | 16 |  |  | 20 |  |  |  | X | X | [13 empty] | 1046 | O/V
25594-1 - Water Treatment System - Reverse Osmosis with Deionizer | Steris | HCS-2052 | 1 | - | - | 66 | 77 | 36 | 1 | X |  | 1548 | Freestanding | FLR | 120 | 1 |  |  |  | 20 | X |  |  | X | [10 empty] | X |  |  | X |  | O/V | X
25825 - Disinfector, Ultrasound Transducer | GE Healthcare (Imaging) | TROPHON EPR | 1 | - | - | 19-1/4 | 13-5/8 | 13-5/8 | 2 | X | 38 |  | WallMount | WAL | 120 |  | 5 | [6 empty] | X | [15 empty] | O/C | X
26370 - Workstation, Prep and Pack, Worksurface 36ft. X 72in. | Steris | ERGOSTAT | 3 | - | - | - | 72 | 45 | 3 |  |  |  | Freestanding | POR | 120 | 1 |  |  |  | 20 |  |  |  | X | X | X | [13 empty] | O/V | X
26691 - Cart, Utility | Intermetro Industries Corp | MW 108 | 2 | - | - | 39 | 24 | 36 | 3 |  |  |  | Mobile | POR | [25 empty] | O/O
27986 - Waste Receptacle, General, 23 GAL, 33"H X 17"W X 20"D | Rubbermaid Commercial | 6146 BEIGE | 16 | - | - | 33 | 17 | 20 |  |  |  |  | Mobile | POR | [25 empty] | O/O
28218 - Cabinet, Flammable Storage, 12 Gal. | Justrite Manufacturing | 891200 | 1 | - | - | 35 | 24 | 18 | 3 |  | 123 |  | Freestanding | [26 empty] | O/O
28657 - Clean Up Counter, 2 Sink | Steris | CCPS317735AH | 1 | - | - | 42 | 77 | 31 | 2 |  | 395 |  | Freestanding |  | 120 | 1 |  |  | 140 |  |  |  |  | X | [5 empty] | X |  |  |  |  | X | X | X | X |  | O/C
28658 - Clean Up Counter, 3 Sink | Steris | CC3110635AH | 2 | - | - | 42 | 106 | 31 | 2 | X | 493 |  | Freestanding |  | 120 | 1 |  |  | 140 |  |  |  |  | X | [5 empty] | X |  |  |  |  | X | X | X | X |  | O/C
28795 - Cabinetry Instrument Storage & Instrument Pegboard | McMurray Stern | Custom | 1 | - | - | 45-3/4 | 238 | 27 |  | X |  |  | Freestanding | VIC | 120 | 1 | [6 empty] | X | [16 empty] | O/O | X
30601 - Air Gun with Wall-Mounting Bracket | Steris | Air Gun with Wall Mounting Bracket | 3 | - | - | - | - | - |  |  |  |  | WallMount | VIC | [15 empty] | X | [9 empty] | O/C | X
Note on V21: Maisie Gwynne: only 1 missing power